#### DATA I/O Corporation

# 10525 Willows Road NE

Redmond, WA 98052

Phone: 1.425.881.6444

Phone: 1.800.247.5700 (press 2)

FAX: 425.867.6972

pmhelp@dataio.com

### Subject: PP100 PC104 Removal and Installation.

Model Affected:  PP100

- PURPOSE:  To inform customers how to change the PC014 on the PP100.

2.0	TOOLS REQUIRED:  Small flat blade screwdriver.

3.0	PARTS REQUIRED:  PC104.

- INSTRUCTIONS:  Turn off the power to PC104 rack on the PP100.  This is done by turning off the circuit breaker on the rear of the PP100.  See figure 1.

Figure 1

  - If you have the old style Autoview you will need to remove the keyboard and monitor cables.  See figure 2.

Figure 2

  - Remove the 4 screws that hold the PC014 in the rack.  2 on the top and 2 on the bottom.  See figure 3.

Figure 3

  - Next remove the PC104 by sliding it out from the rack.  See figure 4.

Figure 4

  - Next slide in the new PC104 into the empty space.  Be careful to align the slides properly.  Do not force the PC104.  It should only take a little pressure to install.  Tighten the 4 mounting screws.

  - Reinstall the cables if you have the old style Autoview.

  - Turn on the power to the PC104 rack.  Programmer circuit breaker.  See figure 1.

  - Go to the PC104 on the display.  Insure that the Sprint menu comes up.  All PC104’s come configured as Programmer 1.  If you are have installed this PC104 on any programmer but 1 you need to edit the Autoexe.bat file.

  - Go to DOS on the PC104.  The edit autoexe.bat.  Edit the line that says Set Station = 1 and change this to the correct Programmer number.  Example Set Station = 3.

# Text boxes

Circuit breaker for Programmers

Remove these 4 screws.